OCT Data Collector　Instruction Manual rev1.1

**1. Start-up**

Run “OctDataCollector.exe”.

OCT Data Collector main screen

**2. Select data**

Select data for output from database or data folder where analysis files are stored (hereafter called as data folder).

**2.1 Select database or data folder**

**In case of database** Choose [Database] and select a database name from pull-down list.

**In case of data folder** Choose [Folder] then click to specify a folder name, otherwise input a folder path directly.

* 1. **Select data**
* Click [Search] to start data search.

Data search window

* Put tick mark on data which you like to output.

Data list window

**＊　Refine your search by capture condition**

　　Search options are shown on [Filter].

|  |  |
| --- | --- |
|  | Item |
| Scan Mode | Circle Scan |
| Radial Scan |
| 3D Scan |
| 3D(V) Scan |
| 3D Wide Scan |
| Include Anterior |
| Fixation  | Disc |
| Macula |
| Center |
| External |
| Wide |
| Eye | Right |
| Left |

**※In case of searching data within selected period.**

In [Searching by term], tick in [Setting a period], specify the desired period, and click [Search].

(In order not to search by term, tick in [All] and click [Search].)

* Click [OK] to finish data selection.

※Selected data are listed in [Export Data List].

　　(“Data File Count” shows the number of selected data. ）

1. **File Export setting**
* Input in [Save Folder] the path to the folder to export files, or browse the folder by clicking .

* Input the prefix of exported file name.

(If nothing is input, “DEFAULT” would be the prefix.)

* Tick in check box in [Patient Info.] to select the patient information exported with data.



|  |  |
| --- | --- |
| item | details |
| Name | Name of patient （Last Name and First Name) |
| Gender | Male or Female  |
| Birth Date | Date of birth |
| Ethnicity | Asian, Caucasian, etc  |
| Ocular Param. | Ocular parameter(Spherical value, Cylinder value, corneal curvature, axial length) |

* Select the protocol and contents of data to export.



|  |  |  |  |
| --- | --- | --- | --- |
| Item | Data to Export | Layer | Setting Detail |
| 3D (Disc) | Disc　3D Scan  | NFL | With Disc Topo Parameter | To export with Disc analysis |
| 3D (Macula) | Macular　3D Scan  | Retina | ETDRS | To export average thickness of each ETDRS section |
| 6 x 6 Grid | To export average thickness of 6x6 grids |
| Circle | Circle Scan | NFL |  |  |
| Radial | Radial Scan | Retina |  |  |
| 3D(V) | Macular　3D(V) Scan | RNFLGCL+GCL++Retina | 10x10　Grid | To export average value of 10x10 Grid in selected layer |
| 6x6 Grid | To export average value of 6x 6 Gred in Retina thickness |
| ETDRS | To export average thickness of each retina thickness’s ETDRS |
| 3D Wide | Wide Scan | RNFLGCL+GCL++Retina | 10x10　Grid | To export average value of 10x 10 Grid in selected layer |
| ETDRS | To export average thickness of each retina thickness’s ETDRS |
| Anterior Radial | Anterior Segment Radial Scan | Cornea |  |  |

※Please refer APPENDIX 1 for details of exported contents.

※Please refer APPENDIX 2 for condition of each exported data.

1. **Export**
* Click [export] to start exporting.

※While exporting, the progress bar is displayed as below. If you cancel exporting, click [Cancel].

※In case when files with the same file name exist in the folder to export, its file name is listed as follows.

\* In order to write over the current data…… select “YES”. Then the current data in the folder is deleted, and start exporting.

\* In order to keep the current data…… select “NO”, and then exporting is canceled. Try to export it again after copying the file to another folder, or changing [prefix].

* When exporting has done, the message saying ”The data output has been completed” is displayed.
1. **Exit the Application**
* Click [Close]

※ In order to save the current setting, check the box of [Save Setting Parameters] and then close the application.

**< APPENDIX 1 > contents of the exported file.**

**◆Additional information（patient / capture information） … common info.**

|  |  |
| --- | --- |
| item | Notes |
| Export Software Ver. | Version of the data collector. |
| Export Date | Year / month / day |
| Export Time | Hour / minute /second |
| Data No. | Data file Number. （FDS/FDA file） |
| Patient ID |  |
| Last Name |  |
| First Name |  |
| Gender | Male / Femalen/a　　　　… No data |
| DOB | Birthday（No data shown as ”n/a”） |
| Ethnicity |  |
| Eye | R ... OUL 　... OS |
| Capture Date | Year / month / day |
| Capture Time | Hour / minute /second |
| Scan Size |  |
| Scan Resolution |  |
| Fixation | Disc / Macula / Center / wideExternal fixation |
| Image Quality |  |
| OCT Focus Mode | VetriousChoroidalCorneaDeep pos |
| Model Name |  |
| Capture Software Ver. |  |
| Analysis Software Ver. |  |
| Analysis Mode | FineBasic |
| Sph. Power | unit: Diopter |
| Cyl. Power | unit: Diopter |
| Corneal Radius | unit: mm |
| Axial Length | unit: mm |
| Contents | RNFL / GCL+ / GCL++ / Retina / Cornea |

**◆3D (Disc)**

File name：XXX\_3D\_DISC.csv

Exported scan size: 6.0 x 6.0 mm / 4.5 x 4.5 mm

|  |  |  |
| --- | --- | --- |
| item | Note | Unit |
| Total | Average thickness of circular | um |
| 4\_T | Average thickness of each 4sector | um |
| 4\_S |
| 4\_N |
| 4\_I |
| 12\_T | Average thickness of each12sector | um |
| 12\_TS |
| 12\_ST |
| 12\_S |
| 12\_SN |
| 12\_NS |
| 12\_N |
| 12\_NI |
| 12\_IN |
| 12\_I |
| 12\_IT |
| 12\_TI |
| Disc Area\* | Disc area | mm2 |
| Cup Area\* | Cup area | mm2 |
| Rim Area\* | Rim area | mm2 |
| Cup Volume\* | Cup volume | mm3 |
| Rim Volume\* | Rim volume | mm3 |
| C/D Area Ratio\* | C/D area ratio |  |
| Linear C/D Ratio\* | Square root of C/D area  |  |
| Vertical C/D Ratio\* | Vertical C/D ratio |  |
| Disc Dia.(V) \* | Vertical disc diameter | mm |
| Disc Dia.(H) \* | Horizontal disc diameter | mm |

\*) Disc topography is applied to “3D Disc, size6.0 x 6.0, resolution 512 x 128”.

12sectors

4sectors

**S**

**SN**

**NS**

**NI**

**IN**

**I**

**IT**

**ST**

**TI**

**TS**

**N**

**T**

**I**

**N**

**T**

**S**

\*This chart is OS. At the OD these grids are mirrored.

**◆3D (Macula) - ETDRS**

File name: XXX\_3D\_MACULA(ETDRS).csv

Exported scan size: 6.0 x 6.0 mm

|  |  |  |
| --- | --- | --- |
| item | Note | Unit |
| ETDRS\_Center | Average thickness of each grid. | um |
| ETDRS\_In\_T |
| ETDRS\_In\_S |
| ETDRS\_In\_N |
| ETDRS\_In\_I |
| ETDRS\_Out\_T |
| ETDRS\_Out\_S |
| ETDRS\_Out\_N |
| ETDRS\_Out\_I |
| Average Thickness | Average thickness of total grid area. | um |
| Center Thickness | The point of ETDRS grids center.（thickness of the center point) | um |
| Total Volume |  | mm3 |

\*This chart is OS. At the OD these grids are mirrored.

**Out\_I**

**Out\_N**

**In\_N**

**Center**

**In\_I**

**In\_S**

**In\_T**

**Out\_T**

**Out\_S**

**◆3D (Macula) – 6 x 6 Grid**

Filename: XXX\_3D\_MACULA (6x6GRID).csv

All 3D macula scan

|  |  |  |
| --- | --- | --- |
| item | note | unit |
| 01\_01 | Average thickness of the each grid. (6 x 6 grid)Each grid named as below. | um |
| 02\_01 |
| 03\_01 |
| 04\_01 |
| 05\_01 |
| 06\_01 |
| 01\_02 |
| 02\_02 |
| ・ |
| ・ |
| ・ |
| ・ |
| 05\_06 |
| 06\_06 |
| Total Volume |  | mm3 |

**6\_1**

**6\_6**

**1\_1**

**1\_6**

**1\_1**

**6\_6**

**6\_1**

**1\_6**

OD

OS

**◆Circle**

File name: XXX\_CIRCLE\_DISC.csv

|  |  |  |
| --- | --- | --- |
| item | Note | Unit |
| Total | Average thickness of circular | um |
| 4\_T | Average thickness of each 4sector. | um |
| 4\_S |
| 4\_N |
| 4\_I |
| 12\_T | Average thickness of each 12sector. | um |
| 12\_TS |
| 12\_ST |
| 12\_S |
| 12\_SN |
| 12\_NS |
| 12\_N |
| 12\_NI |
| 12\_IN |
| 12\_I |
| 12\_IT |
| 12\_TI |

※The name of each sector are same as “3D Disc”.

**◆Radial**

File name： XXX\_RADIAL\_MACULA.csv

|  |  |  |
| --- | --- | --- |
| item | Note | Unit |
| ETDRS\_Center | Average thickness of each grid. | um |
| ETDRS\_In\_T |
| ETDRS\_In\_S |
| ETDRS\_In\_N |
| ETDRS\_In\_I |
| ETDRS\_Out\_T |
| ETDRS\_Out\_S |
| ETDRS\_Out\_N |
| ETDRS\_Out\_I |
| Average Thickness | E Average thickness of total grid area. | um |

※The name of each sector are same as “3D(Macula) – ETDRS”.

**◆3D(V)** - **10x10Grid**

Filename：XXX\_3DV\_MACULA(10x10GRID).csv

|  |  |  |
| --- | --- | --- |
| item | Note | Unit |
| Total | Average thickness of circular | um |
| Superior |  | um |
| Inferior |  | um |
| 01\_01 | Average thickness of each grid. | um |
| 02\_01 |
| 03\_01 |
| 04\_01 |
| 05\_01 |
| 06\_01 |
| 07\_01 |
| 08\_01 |
| 09\_01 |
| 10\_01 |
| 01\_02 |
| 02\_02 |
| 03\_02 |
| 04\_02 |
| 　 |
| 　 |
| 　 |
| 　 |
| 　 |
| 08\_10 |
| 09\_10 |
| 10\_10 |

**10\_1**

**1\_1**

**2\_1**

**1\_2**

**2\_1**

**1\_1**

**10\_1**

**10\_2**



**10\_2**

**1\_2**

**10\_10**

**1\_10**

**10\_10**

**1\_10**

OS

OD

**◆3D(V)** **- 6x6Grid**

File name：　XXX\_3DV\_MACULA(6x6GRID).csv

|  |  |  |
| --- | --- | --- |
| item | Note | Unit |
| 01\_01 | Average thickness of each grid. | um |
| 02\_01 |
| 03\_01 |
| 04\_01 |
| 05\_01 |
| 06\_01 |
| 01\_02 |
| 02\_02 |
| ・ |
| ・ |
| ・ |
| ・ |
| 05\_06 |
| 06\_06 |
| Total Volume |  | mm3 |

※The name of each sector are same as “3D(Macula) – Grid”

**◆3D(V) - ETDRS**

Filename: XXX\_3DV\_MACULA(ETDRS).csv

|  |  |  |
| --- | --- | --- |
| item | Note | Unit |
| ETDRS\_Center | Average thickness of each ETDRS grid. | um |
| ETDRS\_In\_T |
| ETDRS\_In\_S |
| ETDRS\_In\_N |
| ETDRS\_In\_I |
| ETDRS\_Out\_T |
| ETDRS\_Out\_S |
| ETDRS\_Out\_N |
| ETDRS\_Out\_I |
| Average Thickness | Average thickness of total grid area. | um |
| Center Thickness | The point of ETDRS grids center.（thickness of the center point) | um |
| Total Volume |  | mm3 |

※The name of each sector are same as “3D(Macula) – ETDRS”.

**◆3D Wide - 10x10Grid**

ファイル名： XXX\_3D\_WIDE(10x10GRID).csv

|  |  |  |
| --- | --- | --- |
| item | Note | Unit |
| 01\_01 | Average thickness of each grid. | um |
| 02\_01 |
| 03\_01 |
| 04\_01 |
| 05\_01 |
| 06\_01 |
| 07\_01 |
| 08\_01 |
| 09\_01 |
| 10\_01 |
| 01\_02 |
| 02\_02 |
| 03\_02 |
| 04\_02 |
| 　・ |
| 　・ |
| 　・ |
| 　・ |
| 　・ |
| 08\_10 |
| 09\_10 |
| 10\_10 |

※The name of each sector are same as “3D(V) – 10x10Grid”

**◆3D Wide - ETDRS**

Filename:XXX\_3D\_WIDE(ETDRS).csv

|  |  |  |
| --- | --- | --- |
| item | Note | Unit |
| ETDRS\_Center | Average thickness of each ETDRS grid. | um |
| ETDRS\_In\_T |
| ETDRS\_In\_S |
| ETDRS\_In\_N |
| ETDRS\_In\_I |
| ETDRS\_Out\_T |
| ETDRS\_Out\_S |
| ETDRS\_Out\_N |
| ETDRS\_Out\_I |

※The name of each sector are same as “3D(Macula) – ETDRS”.

**◆Anterior Radial**

File name：XXX\_RADIAL\_ANTERIOR.csv

|  |  |  |
| --- | --- | --- |
| item | Note | Unit |
| Center Thickness | thickness of the center point | um |
| Center Curve Radius(H) | Horizontal curve | mm |
| Center Curve Radius(V) | Vertical curve | mm |

**< APPENDIX 2 > Exported data**

Data collector can export the csv-data of scan as below.

◆3D(Disc)

|  |  |  |
| --- | --- | --- |
| Scan protocol | Scan size | fixation |
| 3D | 6.0x6.0mm4.5x4.5mm | DiscExternal |

◆3D(Macula) – ETDRS

|  |  |  |
| --- | --- | --- |
| Scan protocol | Scan size | fixation |
| 3D | 6.0x6.0mm | MaculaCenterExternal |

◆3D(Macula) – 6x6 Grid

|  |  |  |
| --- | --- | --- |
| Scan protocol | Scan size | fixation |
| 3D | 8.2x3.0mm6.0x6.0mm4.5x4.5mm3.0x3.0mm | MaculaCenterExternal |

◆Circle

|  |  |  |
| --- | --- | --- |
| Scan protocol | Scan size | fixation |
| Circle | φ3.4mm | DiscExternal |

◆Radial

|  |  |  |
| --- | --- | --- |
| Scan protocol | Scan size | fixation |
| Radial | φ6.0mm | MaculaCenterDiscExternal |

◆3D(V) – 10x10 Grid

◆3D(V) – 6x6 Grid

◆3D(V) – ETDRS

|  |  |  |
| --- | --- | --- |
| Scan protocol | Scan size | fixation |
| 3D(V) | 7.0x7.0mm | MaculaExternal |

◆3D(V) – 10x10 Grid

◆3D(V) – ETDRS

|  |  |  |
| --- | --- | --- |
| Scan protocol | Scan size | fixation |
| 3D Wide | 12.0x9.0mm | WideExternal |

◆Radial Anterior

|  |  |  |
| --- | --- | --- |
| Scan protocol | Scan size | fixation |
| Radial | 6.0x6.0mm | External |

End.